## Slide 1: Device programming for CRT optimization
David Luria, MD
Hadassah Medical Center

## Slide 2: Conflict of interest
NONE

## Slide 3: “What is your routine practice for CRT optimization?”
”Im not sure I remember what is it!”

## Slide 4: Non responders- continue challenge

## Slide 5: Optimization: concept
Patient selection
Medical therapy correction
Guarantee 100% pacing (to treat PVC, AF)
Intraoperative optimization (lead position )
Programming optimization (AV/VV, MPP)

## Slide 6: Electrical or mechanical resynchronization?
Electrical Concept: “Electrical therapy for electric problem”
Solution: pace at the latest activation (q-LV interval) and target shortest pace QRS
Mechanical concept: Final result is contraction improve
Solution: use ECHO, dP/dt guidance and target best hemodynamic parameters

## Slide 7: Lead placement
Lateral wall (postero-lateral, antero-lateral)
Non apical (MADIT CRT, REVERCE )
Latest activation (Q-LV or RV-LV)
Latest contraction (ECHO, speckle tracking)
Escape scared myocardium (ECHO, radial strain)
Multisite/multipolar pacing

## Slide 8
23 consecutive pts
All available veins
Acute electrical & hemodynamic measurements
Q-LV = 95 msec ensure > 10% increase of dP/dt in all pts

## Slide 9: Q-LV & RV-LV intervals for prediction of CRT response

## Slide 10: Speckle Tracking Radial Strain Imaging - guided Lead Placement For Improving Response to CRT In Patients With Ischemic CMP - The Raise CRT Trial -
Michael Glikson(1), Gregory Golovchiner (2) , Moshe Swissa (3), Monther Boulus (4) Sami Viskin (5) Aharon Medina (1), Moti Haim (6) , Paul Friedman (7) , Vladimir Khalameizer (8), Ori Vatury (9 ) , Ito Saki (7), Nir Shlomo (10), Jae K. Oh (7), Ilan Goldenberg (10), Roy Beinart (9)

## Slide 11: Combined outcome 12 m – MACE (death, CHF hospitalization )

## Slide 12
Results:
  20% procedure/system related complications
  At 12 and 24 months no difference in:
  - Hospitalizations
  - NYHA
  - EF

## Slide 13: Lead’s advances
Quadripolar leads
Passive fixation
Active fixation

## Slide 14
44 pts
1:1 randomization for MPP
12 months f/u

## Slide 15
110 consecutive pts
Intraoperative lead position optimization by Q-LV and
dP/dt max
1 year follow up
“RESPONDER”
  - 15% ESV
  - I NYHA
  Packer’s score =0

## Slide 17: Timing programing optimization
AV delay – ventricular filling optimization
  Too long (early atrial contraction)– decrease atrial diastolic feeling
  Too short – decrease atrial kick
VV delay – synchronization optimization

## Slide 19: Optimization: limitations
Not adapting to changing conditions:
  Orthostasis
  Exercise
  Loading conditions
  Drug regimen
Time and resource consuming
Clinical significance of observing hemodynamic changes are questionable

## Slide 20
Prospective randomized 1:1:1
  FIXED AV delay = 120 ms
  ECHO optimization by mitral inflow
  BOSTON SC SmartDelayTM
6 month clinical/ECHO follow up
NO DIFFERENCE in any parameter

## Slide 21: GUIDELINES citations

## Slide 22: Device-based optimization algorithms

## Slide 23: BOSTON SC – SmartDelayTM
Target: to allowed native septum activation
Method: paced/sensed AV and RV/LV conductions measurements (AV optimization only)
Programming: run of semi-automatic algorithm after implant
NOTE: frequent only LV pacing and long AV delay
DATA:
  Favorite small hemodynamic studies
  SMART AV prospective randomized clinical trial – NEGATIVE results
  Some advantage in subgroups of pts

## Slide 24
SMART AV study sub-analysis
280 pts
Randomization
  Fixed AV 120
  SD (EG based algorithm)
End point: ESV reduction >15%

## Slide 25: MEDTRONIC: AdaptiveCRT TM BIOTRONIC : AutoAdaptTM
Target: to allowed native septum activation
Method: Dynamic every minute AV measurement
Programming: automatic AV/VV optimization

## Slide 26
Prospecting randomized trial
2:1 aCRT to ECHO optimization
6 month follow up
Non-inferiority primary end points
  Clinical score (Packer)
  Echo AoVTI

## Slide 27: AdaptiveCRT TM
Sub-analysis of “only LV pacing” group from Adaptive CRT trial. HR 2013 ; 10: 1368
44,838 pts from CARE LINK data base Abstract from HRS 2017

## Slide 28: ABBOT (SJM): Sync AV TM
Target: AV pacing by programmed offset
Method: Dynamic (every 256 beats) intrinsic AV measurement
Programming: automatic. AV offset programmable (-50 to -10)
Retrospective
Propensity matched
4:1
Abstarct, HRS 2019

## Slide 29: CONCUSIONS
There are NO class I/IIA evidence/recommendations for intraoperative or programming optimization
Lead location selection for non apical area with latest
(more than 75-95 msec) Q-LV /RV-LV may be useful
Multipolar pacing from single quadripolar lead is promising
New automatic algorithms of AV/VV optimization demonstrated non inferiority as compared to ECHO guided optimization in randomized studies
LV pacing only/before RV may save buttery life and improve clinical results in subset of patients

## Slide 30: Back up

## Slide 34: Reasons for insufficient response rate
Internal limitations –desynchronization is only part of the problem
  LBBB cardiomyopathy/pacing cardiomyopathy – very high responding rate
  Severe ischemic CMP – relatively low
Patient selection: RBBB vs LBBB, QRS width, AF, pt gender

## Slide 35: Reasons for insufficient response rate
Absence of robust criteria for best lead position
Technical challenges in lead placing
  Limited choice of veins
  Low stability in proximal vein (basis LV)
  High threshold/diaphragmatic pacing
  Long procedure/Xray/contrast
Non physiological pacing
  epicardial
  partially synchronized